## Slide 1
Prompt gamma detectors for
Hadrontherapy and BNCT
Carlo Fiorini
Politecnico di Milano and INFN, Milano, Italy
CNAO-IN2P3 Meeting 05-03-26

## Slide 2: The REALPATH project
REALPATH - A Real-Time Range Verification System for Hadron Therapy
2
Mechanically collimated camera, pushing for a compact detector paylod
High-speed electronics platform to cover different accelerator types
Integration in the treatment room (including software and safety aspects)
Clinical validation (possibly in both fixed-line and gantry-based facilities)

## Slide 3: Idea of inverted collimator
3

## Slide 4: Validation of the inverted collimator
4
Features:
Falloff Retrieval Precision (@108 protons) similar to KE collimator
Much smaller collimator weight (7kg vs. 25kg)

## Slide 5: Camera geometry and expected Falloff Retrieval Precision (FRP)
5
Experimental measurements scalded to the final PGI Camera detection area:
10cm x 20cm

## Slide 6: Detection Module Design
6
Compact 64-channel base unit comprising:
  64 LYSO (Lu1.8Y0.2SiO5) pixelated crystals
  64 SiPMs matrix cell pitch
Front-End PCBs with ASIC read-out and FPGA processing:
  Energy range up to 10 MeV
  Count Rate per channel 1 Mcps
  TDC precision (FWHM) < 250 ps
Flexible Configuration
  Modular, Compact, and Scalable design
  Linear and angled arrangements
  Overall FOV (Field Of View) > 5 cm
Compatible design for future channel expansion: 64ch 256ch 512ch
C. Riboldi et al., " Design of a 64-channel modular platform for high-speed readout of scintillator-based detectors for range monitoring in hadrontherapy“, JINST 2025.

## Slide 7: First camera prototype (under construction)

## Slide 8: First spectroscopy tests
Measurements with a 16-channel detector tile and 22Na source:
excellent energy resolution
  10.7% average FWHM at 0.511 MeV
  6.7% average FWHM at 1.275 MeV
homogeneous resolution of all pixels

## Slide 9: High rate measurements (last week!)
Succesful tests (137Cs source) up to 1Mcps/channel!
662keV

## Slide 10: REALPATH instrument positioning system
10
Fixed-beamline systems
(e.g. CNAO)
Trolley
Suspended
Gantry-based systems
(e.g. IBA systems)
L.Santini, et al., «Towards a novel positioning system for a real-time prompt gamma imaging detector in hadrontherapy», to be submitted to JINST.

## Slide 11: Gamma/Neutron Discrimination by means of CLYC and PSD
Pulse Shape Discrimination (PSD) :
PSD
CLYC scintillator and readout electronics:
O.Halfon, et al., IEEE TNS, VOL. 72, NO. 7, JULY 2025.

## Slide 12: PGI measurements with real-time neutron-gamma discrimination
O.Halfon, et al., “First Experimental Measurement of Prompt Gamma Imaging with Real-Time Neutron Rejection via Pulse Shape Discrimination”, IEEE TRPMS, 2026.

## Slide 13: Boron Neutron Capture Therapy
Goal: Development of a SPECT (Single Photon Emission Thomography) system) for BNCT
Detection of emitted 478keV gamma photons may let to estimate 10B neutron captures and support therapeutic outcome (personalized dosimetry).
Scintillator/SiPMs-based gamma-ray detector
Neural-Network for event reconstruction
BeNEdiCTE (Boron NEutron CapTurE) detector

## Slide 14
Tomographic Measurements
Neutron Flux
T. Ferri et al., “Design and Validation of a SPECT Prototype for Treatment Monitoring in BNCT and First Experimental Tomographic Results,” IEEE Transactions on Radiation and Plasma Medical Sciences.
Reactor power = 70 kW
Measurement time = 36 min
Neutron flux =
0° Projection
120° Projection
60° Projection
180° Projection
The 3D reconstruction was obtained using STIR and 10 iterations of the OSEM algorithm.
LENA
Pavia (Italy)

## Slide 15
Nagoya
University (Japan)
The NUANS (Nagoya University Accelerator-driven Neutron Source) is composed by a dynamitron accelerator, a lithium target assembly and a BSA.
Nishitani, Takeo, et al. "Neutronics Analyses of the Radiation Field at the Accelerator-Based Neutron Source of Nagoya University for the BNCT Study." Journal of Nuclear Engineering 3.3 (2022): 222-232.
Epithermal neutron flux that fulfils IAEA standards;
Measurements with a water-filled phantom;
Boron background due to the walls.
Phantom
Irradiation room
Portland cement
Borated polyethilene
Toward BNCT clinical environment

## Slide 16
Nagoya
University (Japan)
20x20x20 cm3 water-filled phantom with an inserted vial containing different boron concentrations
Neutron flux of approx 1x107n/cm2/s
Shield2
Polyethylene
Phantom
Nozzle
Vial
Cadmium
Lead
Collimator
Epthermal neutrons are thermalised interacting with the shielding and are captured by the 10B inside the system electronics
145 cps
Water
60 cps
Toward BNCT clinical environment
Shield1
Polyethylene
Cadmium

## Slide 17
Nagoya
University (Japan)
Despite the high background, we were able to image vials with varying 10B concentrations down to 1500 ppm
Decreasing 10B concentration
Neutron flux ~ 1x107n/cm2/s
Vial filled with 3686 ppm of 10B
Vial filled with boron powder
~ 2.7 cm
~ 2.7 cm
y
3686 ppm of 10B + boron powder
A. Caracciolo, D. Mazzucconi, T. Ferri, L. Grisoni, F. Ghisio, et al., “Prompt gamma-ray imaging in realistic background conditions of a boron neutron capture therapy facility”, Scientific Reports.
Toward BNCT clinical environment

## Slide 18: Acknowledgments
REALPATH
Politecnico di Milano, DEIB: K.Urban, M.Piroddi, L.Santini, C.Riboldi, A.Forgione, O.Halfon, F.Garzetti, A.Geraci, E.Ronconi, G.Borghi, M.Carminati, G.Baroni
Polimi di Milano, ENG/CeSNEF: E.Bellotti, D.Mazzucconi, D.Bortot, S.Agosteo
CNAO (Pavia, Italy): M.Ciocca, M.Donetti, M.Pullia, A.Pella
BNCT
Politecnico di Milano, DEIB: F.Ghisio, T.Ferri, Pandocchi, G.Borghi, M.Carminati
Polimi di Milano, ENG/CeSNEF: G.Mocerino, L.Grisoni, D.Mazzucconi, S.Agosteo
Università di Pavia: N.Protti, S.Altieri, V.Pascali
Università di Bari: D.Ramos, G.Iaselli, G.Pugliese
Fundings: PRIN-MUR, INFN CSN5