## Slide 1: Matterhorn, a single photon counting detector for high flux synchrotrons applications
Roberto Dinapoli, Aldo Mozzanica, Anna Bergamaschi, Bernd Schmitt, Carlos Lopez Cuenca, Christian Rueder,Davide Mezza, Dhanya Thattil, Dominic Greiffenberg, Erik Fröjdh, Jiaguo Zhang, Jonathan Mulvey, Khalil Daniel Ferjaoui, Konstantinos Moustakas, Mar Carulla Areste, Martin Brückner, Martin Mueller, Saverio Silletta, Shuqui Li, Simon Ebner, Thomas King, Vadym Kedych, Viktoria Hinger, Viveka Gautam, Xiangyu Xie.
13th FEE workshop, Paris France, 18-22 May 2026

## Slide 2: Paul Scherrer Institute
Energy research
Hot lab
Radiopharmaceuticals
Biology
X-ray Free Electron
Laser SwissFEL
Nanotechnology
Neutron source
SINQ
Muon source
SµS
Synchrotron source
SLS
Particle physics
Proton Therapy
Proton
Accelerator
Park Innovaare
Notes: Energy, radio pharma, biology, proton therapy.
-SINQ: neutrons
- Sus: muons
- Protons
- Synchrotron
- XFEL

## Slide 3: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021

## Slide 4: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
For the detector this means:
Faster count rate
Higher frame rate
Wider energy range
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021

## Slide 5: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
For the detector this means:
Faster count rate
Higher frame rate
Wider energy range
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021
- 15-20 M photons/pixel/second
- 20 kHz continuous frame rate
- 250 eV – 80 keV

## Slide 6: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
For the detector this means:
Faster count rate
Higher frame rate
Wider energy range
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021
- 20 kHz continuous frame rate
- 250 eV – 80 keV
LGADs
High Z
- 15-20 M photons/pixel/second

## Slide 7: SLS 2.0 – a 4th generation synchrotron
Increased coherent flux [1]:
  >100x at 10 keV
  Up to 1000x at 20 keV
Increased electron energy (2.4  2.7 GeV)
  Higher energy photons (80 keV)
For the detector this means:
Faster count rate
Higher frame rate
[1] SLS 2.0 Storage Ring Technical Design Report PSI Bericht Nr. 21-02 November 2021
- 20 kHz continuous frame rate
Charge integration
Single Photon Counting
- 15-20 M photons/pixel/second

## Slide 8: (Historical) detector portfolio: Single photon counting
 | MYTHENI-II-III | PILATUSI-II | EIGER | MATTERHORN
Technology | UMC 250/110 nm | DMILL800/UMC 250 nm | UMC 250 nm | UMC 110 nm
Status | Modules available | Modules available | Modules available | Prototyping phase
Channel size | 50 µm (Strips) | 172 x 172 µm2 | 75 x 75 µm2 | 75 x 75 µm2
New

## Slide 9: (Historical) detector portfolio: Single photon counting
 | MYTHEN III | PILATUS | EIGER | MATTERHORN
Technology | UMC 110 nm | UMC 250 nm | UMC 250 nm | UMC 110 nm
Status | Modules available | Modules available | Modules available | Prototyping phase
Channel size | 50 µm (Strips) | 172 x 172 µm2 | 75 x 75 µm2 | 75 x 75 µm2
New
Phased out
(DECRTIS)
Phasing out

## Slide 10: The “speed challenge”: high photon flux
Single photon counting works extremely well if the incoming photon rate is reasonable low

## Slide 11: The “speed challenge”: high photon flux
Single photon counting works extremely well if the incoming photon rate is reasonable low
But as the rate increases we lose efficiency due to pulse pileup

## Slide 12: The “speed challenge”: high photon flux
Single photon counting works extremely well if the incoming photon rate is reasonable low
But as the rate increases we lose efficiency due to pulse pileup
Pile-up mitigation techniques for SPC:
Time over threshold
Retriggering
…

## Slide 13: Pile-up counting with multiple thresholds
Paralyzable counter:
Probablility of two and three events pile-up:
Th as fraction of photon energy
[1] Glenn F. Knoll, Radiation Detector and Measurement 4th Edition, Wiley
[2] M. Andrae The MYTHEN III Detector System - A single photon counting microstrip detector for powder diffraction experiments ETHZ Doctoral Thesis
Counter 0
Counter 1
Counter 2
50%
130%
170%
100%
100%: Energy of the monochromatic incoming photons

## Slide 14: Multi-thresholding : proof of concept with Mythen3.02 (1D)
Settings: standard
Energy: 15 keV
90% efficiency at:
Single counter : 1.03Mcps
MS Beamline @ SLS

## Slide 15: Multi-thresholding : proof of concept with Mythen3.02 (1D)
Settings: standard
Energy: 15 keV
90% efficiency at:
Single counter : 1.03Mcps
Three counters : 6 Mcps
Noise: 175e- RMS
MS Beamline @ SLS

## Slide 16: Multi-thresholding : proof of concept with Mythen3.02 (1D)
Settings: standard
Energy: 15 keV
90% efficiency at:
Single counter : 1.03Mcps
Three counters : 6 Mcps
Noise: 175e- RMS
MS Beamline @ SLS
Settings: Fast [1]
Single counter 3.52 Mcps
Three counters 20.87 Mcps
Noise:~250e-
[1] M Andrä The MYTHEN III Detector System - A single photon-counting microstrip detector for powder diffraction experiments ETHZ PhD Thesis 2021
x6
x6

## Slide 17: Matterhorn – a new single photon counter for SLS2.0
 | EIGER | MATTERHORN
Technology | UMC 250 nm | UMC 110 nm
Pixel size | 75 x 75 um2 | 75 x 75 um2
Thresholds | 1 | 4
Counter depth | 12 bit (3x4) | 4x16 bit
90% efficiency | 350k counts/pixel/s* | 15-20M counts/pixel/s
Module size | 4x8cm2 | 4x8cm2
HighZ/LGAD | No | Yes
*DECTRIS EIGER2 90% efficiency at ~2 Mcps (UMC 110nm)
T. Donath et. al. EIGER2 hybrid-photon-counting X-ray detectors for advanced synchrotron diffraction experiments J. Synchrotron Rad. (2023). 30, 723-738

## Slide 18: Matterhorn 0.1, 12.2022
ASICs: ~5x5 mm2 (3.6x3.6 active=48x48=2304 pixels)
MPW chips bump-bonded to sensor (XR-Nanotech)
Matterhorn 0.2,10.2023
One pixel version only
2x 3.125 GHz serializers
3.125 GHz CML driver
Internal DACs for biasing
Some “features” fixed
Matterhorn prototypes
2 Pre/shaper variations
2 Comparator variations
2x 1.6 GHz serializers
1.6 GHz CML driver (with pre-emphasis)
Analog variations
“Slow” readout
One version only
Fast readout

## Slide 19: Matterhorn 0.3, 12.2024 new digital periphery features
Matterhorn prototypes

## Slide 20: Matterhorn 0.3, 12.2024 new digital periphery features
Matterhorn prototypes
Matterhorn 1.0,t.o. 08.2025 Delivered: 01.2026
MH0.2
MH 1.0
~2 cm

## Slide 21: Matterhorn 1, the full-size chip
Pixel data transfer:
  Current mode signal distribution (to pixel periphery)
New periphery, additional features:
  Global signals distribution (RES,STO,EN<0:3>)
  Counter summing mode
  4 bit mode
  64b/66b encoder for serializers
  SPI interface for configuration/slow control
Gating (4 independent EN) <20ns
No active circuitry on the chip sides (less dead area)
 | MATTERHORN 1
Technology | 110 nm
Chip active area | 1.92x1.92 cm2 256x256 pixels
Pixel size | 75 x 75 um2
Thresholds | 4
Counter depth | 4x16 bit
Trim bits/compar. | 6
Static power/pixel (4 comparators) | 9.6μW
HighZ/LGAD | Yes
Same pixel as Matterhorn 0.2
New features
MH0.2
MH 1.0
~2 cm

## Slide 22: Matterhorn pixel architecture
Pixel
1)
1) Both e- and h+ readout possible
1)
1)
(Vrf voltage)
(Vrfsh voltage)
Static power/pixel (4 comparators) | 9.6μW

## Slide 23: Matterhorn pixel architecture
1)
1) Both e- and h+ readout possible
1)
1)
(Vrf voltage)
(Vrfsh voltage)
Static power/pixel (4 comparators) | 9.6μW
We had problems in UMC110 with small MIMCAPS
We moved to MOMCAPS where possible

## Slide 24: Pixel testability, charge injection
5)
1)
(Vrf voltage)
(Vrfsh voltage)
Cinj
BOC pulser
OLD, MH0.2
5) The preamp can be pulsed with a programmable voltage step
5)

## Slide 25: Pixel testability, charge injection
1)
1)
1)
(Vrf voltage)
(Vrfsh voltage)
Cinj
PIX pulser
BOC control
NEW, MH1.0
5) The preamp can be pulsed with a programmable voltage step
6)
6) Each counter can individually be pulsed

## Slide 26: Pixel testability, analog out
Differences wrt EIGER:
No per-pixel analog out
Very good buffers
Pre
Shaper
Comp2
Count2
(255,255)
Pre
Shaper
Comp2
Count2

## Slide 27: Matterhorn pixel architecture
Pixel
2)
2) One comparator can be connected to two counters (pump-probe)

## Slide 28: Matterhorn pixel architecture
Pixel
3)
3) Each comparator can individually be ‘masked’ (disconnected from the counter)
 reduce switching for noisy channels

## Slide 29: Matterhorn pixel architecture
Pixel
4)
4) Each comparator can individually be powered-off when noisy or not used

## Slide 30: Matterhorn counters
T
D
T
STORE
D
STORE
Counter T- flip-flops
RESET
RESET
Enable
D- latches for continuous R/W
Very dense full custom layout with 0.87 transistors/um2 (physical dimensions)
(EIGER: storage on MIMCAPs to save silicon area)
All logic on Triple Well to reduce D/A crosstalk

## Slide 31: Matterhorn 1.0 counters readout
D
STORE
D- latches for continuous R/W
ROW SELECT
Column Data Bus
16 lines
1024x Current comparators
Chip digital periphery
MH0.2: Voltage readout
MH1.0: current readout
Same concept as EIGER but:
Tunable Ipixel
New current comparator
Ipixel

## Slide 32: Periphery blocks
PLL
  Locking range: 930 MHz ÷ 2600 MHz
  Power: 18 mW
  Jitter < 5 ps
All biases available on internal DACs (or self bias)
Transistor threshold monitor
Ring oscillator
Some critical biases also available from on-board DACs
4 separated comparator biases

## Slide 33: Periphery blocks
Pixel & counter selector
Counter summing mode
SPI interface for configuration/slow control
RESET, STORE flexible global distribution
Very high speed ENABLE (gating) distribution
CML driver
  Main driver and pre-emphasis current tunable (3 bits): 3 mA ÷ 15 mA
  Pre-emphasis delay tunable (2 bits)
Pixel/counter readout logic (includes 16,8, 4 bit mode)
64/66b Aurora encoding
4xHigh speed serializers @3.125MHz
NO logic on the chip sides
How many columns at the same time and delay are selectable
Targeting 20ns distribution delay for pump-probe/hybrid SLS mode
All periphery logic synthesized with digital tools for the first time!

## Slide 34: Single pixel S-curves
All counters enabled but cnt1,2,3 at 4095 DAC
All counters scanned
Counter0
Default settings
(Vrf=990DAC#)
Cu fluorescence
MH1.0

## Slide 35: Threshold energy calibration
All counters scanned
Counter0
Default settings
(Vrf=990DAC#)
Multiple fluorescence targets
MH1.0

## Slide 36: Threshold dispersion at 9 keV
vrf | Untrimmed [eV] | Trimmed [eV]
200 | 998 | 59
300 | 905 | 43
400 | 668 | 33
500 | 477 | 29
Small chip (48x48 pixels)
Non optimized trimming procedure
Trimming on inflection point
8e-
:16 to 20
MH0.2
130e-

## Slide 37: Threshold distribution
Sigma before trimming:
~130e-
Default settings
Low- noise
settings
MH1.0
Defective columns

## Slide 38: Noise performance
Counter0
Default settings
(but Vrf scanned)
Cu+Mo fluorescence
Copper fluorescence
Molybdenum fluorescence
MH1.0
DAC units are different in MH1.0 and MH0.2
fast - slow

## Slide 39: Effect of Icomp on detected incoming rate
Icomp Default (1200 DAC)
Icomp 3x (2495 DAC)
Counter0
Default settings
(but Vrf scanned)
Beamtime at PSI
MH0.2
1.3x
1.45x
1.4x
τ=18.7ns
Single counter

## Slide 40
Vicomp 2495 DAC (3x current)
Vrf 200 DAC
Multi-threshold rate tests at PSI with MH02
MH0.2
E=12 keV
Fits: paralyzable model
Vicomp 1200 DAC (1x current)
Vrf 200 DAC
Fits: paralyzable model
x2.1
x3.1

## Slide 41
Vicomp 2495 DAC (3x current)
Vrf 300 DAC
Multi-threshold rate tests at PSI with MH02
MH0.2
E=12 keV
Fits: paralyzable model
Vicomp 1200 DAC (1x current)
Vrf 300 DAC
Fits: paralyzable model
x2.7
x3.6

## Slide 42
Vicomp 2495 DAC (3x current)
Vrf 400 DAC
Multi-threshold rate tests at PSI with MH02
MH0.2
E=12 keV
Fit: paral. model
Fit: sum model
Fit: paral. model
Fit: paral. model
Vicomp 1200 DAC (1x current)
Vrf 400 DAC
x3.1
x6.9

## Slide 43: The “speed challenge”: high frame rate
 | EIGER | MATTERHORN 1
Technology | 250 nm | 110 nm
Chip outputs | 32xCMOS, 200 Mb/s | 4xCML, 3.125 Gb/s
Single chip | 6.4 Gb/s | 12.5 Gb/s
Module (8 chips) | 51.2 Gb/s | 100 Gb/s
Readout channels/module | 2x10Gbit/s | 1x100Gbit/s
Frame rate | ~10 kHz 8bit (burst) | ~20 kHz 8bit (continuous)
Continuous R/W | Yes | Yes

## Slide 44: Summary
Matterhorn is a new SPC detector for SLS2.0 targeting:
  15-20Mcps count rate with 10% loss of efficiency
  250 eV – 80 keV energy range (with LGADs and HighZ)
  20kHz frame rate; <20 ns gating
Results from MH01, MH02 and full-size chip MH1.0:
Noise down to <~80e- in slow settings
Demonstrated ~14.4 Mcps at 90% efficiency using multiple thresholds with MH0.2
Tau <20ns in single counter mode
Digital periphery and high speed links working
MH0.2
MH 1.0
~2 cm

## Slide 45: Outlook
Detailed characterization of MH1.0
  Rate tests
Parameters optimization
  Maximize speed with pile-up counting
I/O speed
Radiation tolerance
High-Z and LGAD integration
Full detector development
  Module Q1 2027?
JUNGFRAU detectors of different size
500k
1.5M
4M
16M

## Slide 46
THANKS!

## Slide 47
V. Kedych, M. Müller, P. Sieberer, K. Ferjaoui, M. Brückner, E. Fröjdh, J. Heymes, C. Lopez-Cuenca,
K. Moustakas, C. Ruder, J. Mulvey, S. Silletta, V. Hinger, A. Mozzanica, M. Carulla, X. Xie,
R. Dinapoli, A. Bergamaschi, S. Li, D. Thattil, J. Zhang, A. Mazzoleni, V. Gautam, D. Mezza, T. King, B. Schmitt.
Missing: S. Ebner, D. Greiffenberg, S. Hasanaj, S. Emiliani, L. Lunid, S. Moses, M. Anastasi, I. Giannopoulos, M. Huerst
THANKS!

## Slide 48
Backup slides

## Slide 49: Pile-up mitigation techniques
Use charge integration with dynamic gain switching (Jungfrau, Jungfrau 2)
Time over threshold/retriggering
Credit: R. Ballabriga

## Slide 50: Operating modes of a single photon “counter”

## Slide 51: Comparison of pileup mitigation
[1] A. Bergamaschi et. al. Time-over-threshold readout to enhance the high flux capabilities of single-photon-counting detectors J. Synchrotron Rad. 18, 923-929. 2011
[2] P. Zambon, Dead time model for X-ray photon counting detectors with retrigger capability, NIMA 2021

## Slide 52: Pixel testability
Pixel
5)
5) The preamp can be pulsed with a programmable voltage step
Preamp and shaper outputs of pixel 0,0 can be monitored
6)
6)
6)
6)
6) Each counter can individually be pulsed
Comparator and “count” outputs of pixel 0,47 can be monitored
Pre
Shaper
Comp2
Count2

## Slide 53: The «energy challenge»: low photon energies
  Low quantum efficiency (QE):
  Development of “thin entrance window” sensor
  Low signal-to-noise ratio (SNR):
        Increase the signal with inverse Low Gain Avalanche Diodes (iLGADs) with a multiplication factor 5-20
  Results: 80 % QE for a planar pin diode @250eV
  Results: nominal noise of 35 e- with standard sensor at RT down to 3.7 e- at -22 oC with Moench
Soft X-ray applications at SwissFEL and SLS (250 eV – 2 keV)

## Slide 54: The «energy challenge»: high photon energies
CdZnTe is the material of choice above 15 keV
4×4 cm² sensor commercially available
Dynamic behavior (signal stability, afterglow) close to silicon sensors
Higher photon energies need to be studied
Crystal defects influence sensor uniformity
Fluorescence photons with long absorption range
GaAs:Cr is promising, but doesn‘t perform as well as CdZnTe
Two European suppliers, 8×4 cm² sensors available
 Afterglow acceptable up to photon fluxes
 Continuous: <10(8-9) ph/mm²·s  Pulsed: <10(6-7) ph/mm² ·pulse
Some signal loss at photon fluxes above > 3×1010 ph/mm²·s
Crystal defects influence sensor uniformity, variation of effective pixel size + Short hole lifetime  Craters
High dark current  For long exposure times: ideally SPC
      Hard x-ray applications at Synchrotrons, 15-80keV

## Slide 55: Timeline
First prototype back!
First prototype submitted (MPW)
Second prototype (MPW) Submitted
SLS2.0
Matterhorn
Second prototype back!
Single module detectors?
Full ASIC submission
Multi module
detectors
Planned submission
of third prototype
Move to new building

## Slide 57
Default settings:
For Mh02 , external DACs: 400 DAC. For Mh10, internal DACs: 990 DAC

## Slide 58: Basic characterization of MH02
Single pixel data
Metrologie beamline at SOLEIL with help from: Marie Andrä and Arkadiusz Dawiec
Calibration at: 7, 9, 12, 15 keV
Threshold scans fitted with analytical model:

## Slide 59: Four counters scan
Scanned at the same time
MH 0.2

## Slide 60: Noise performance of MH02
fast - slow
Preamplifier feedback voltage
(gain)

## Slide 61: Single counter rate tests at MS beamline - MH01
Predicted rate capability using analytical model
Measurement

## Slide 62: Single counter rate tests at MS beamline – MH01
Note: slightly over estimated due to kB peak

## Slide 63: Single counter rate tests at MS beamline – MH01
Noise vs. rate capability
E=12 keV

## Slide 64: Multi threshold rate tests at SOLEIL with MH02
Sum counter 0-3
Counter 0
x6

## Slide 65: Multi threshold rate tests at SOLEIL with MH02
Sum counter 0-3
Counter 0
Note: fitted with simple paralyzable function
Sum counter 0-3
Counter 0
x6
x3

## Slide 66: Multi threshold rate tests at SOLEIL with MH02
Sum counter 0-3
Counter 0
Sum counter 0-3
Counter 0
x6
x3
Comparator biased in “low power” mode
Vrf can be set even lower (100-150)
Preamp/shaper biases can be further tuned

## Slide 67: Trimming with noise (very preliminary)
Noise peak, NOT threshold distribution
Noise peak moves quite linearly with TB value
counts
Default settings
(Vrf 990 DAC#)
MH1.0

## Slide 68
Trimming with noise (very preliminary)
Noise peak, NOT threshold distribution
~110e-
~12e-
:9
counts
Default settings
(Vrf 990 DAC#)
130/9<15e-
MH1.0
Full chip (256x256 pixels)
Non optimized trimming procedure
Trimming on noise