## Slide 1
XIDyn
High Flux, High Dynamic-range Hard X-ray detector full-reticle ASIC for 4th generation synchrotrons
Luke Mallett (RAL STFC)
On behalf of the XIDyn Collaboration

## Slide 2
Introduction
New generation of synchrotrons operate at significantly higher flux and higher energy photons
Faster high flux detector systems required
Must use higher Z materials (CZT)
ESRF
Diamond II
Up to photons/s/

## Slide 3
Specification
Parameter | Continuous | Burst Timing
Operation | Constant with tolerable deadtimes | 256 Frame capture and slower readout
Master Clock Frequency [MHz] | Up to 684.5 | 500
Energy Range [keV] | 25keV typical | 30-100keV
Resolution [photon] | 0.25-1 | 1
Noise [e] | <1200 (1/4 photon) | <1200 (1/4 photon)
Flux [ph/pixel/s] | ~1010 | -
Pixel Size [µm] | 110 | 110
Frame Rate [kHz] | ~10-150 | 5700 for 16 bunch mode
Instantaneous Flux [photons] | ~100 | ~200
Power [W/cm2] | 5 | 5
Notes: 110um pitch – compromise between resolution and data output

## Slide 4
Introduction
Rutherford Appleton Laboratory developed DynamiX
University of Heidelberg developed XIDer
Lots of similarities:
  65nm process
  Hybrid design for CZT detector
  Aiming for photons/s/pixel
  16x16 pixels
Test structures shared an MPW
BabyDynamiX
XIDer T6
2mm Redlen HF-CZT
Both shown at
FEE 2023
Notes: BOTH PRESENTED AT FEE 2023

## Slide 5
Introduction
XIDyn
Efforts combined to XIDyn (XIDer + DynamiX)
192x144 pixels reticle scale ASIC
Pads only on one edge for tiling
Using RAM in pixel from XIDER
Maximum of 20W
21mm
17mm
Notes: 144 rows as 192 rows too large – IR drop/timing

## Slide 6
ASIC Architecture
Large pixel array controlled by global digital logic
3 Dual lane 14.1Gbps serializers
Using 4x4 super pixels
Digital-on-top approach
Design challenges:
  IR drop up columns
  Pixel mismatch
GLOBAL DIGITAL LOGIC | FIFO | GLOBAL DIGITAL LOGIC | FIFO | GLOBAL DIGITAL LOGIC | FIFO | GLOBAL DIGITAL LOGIC
 | SERIALISER |  | SERIALISER |  | SERIALISER | 
PADS
PIXEL ARRAY

## Slide 7
ASIC Architecture
16 Analogue Channels
Digital logic
16 RAMs
(256x16b)
GLOBAL DIGITAL LOGIC | FIFO | GLOBAL DIGITAL LOGIC | FIFO | GLOBAL DIGITAL LOGIC | FIFO | GLOBAL DIGITAL LOGIC
 | SERIALISER |  | SERIALISER |  | SERIALISER | 
PADS
PIXEL ARRAY

## Slide 8
Analogue Channel
Charge cancellation
Charge is integrated and removed when amount exceeds threshold
Allows for constant charge removal from Cf to prevent saturation
Removes charge in known packet sizes and directly provides digital output (no ADC)
With the high flux, cannot achieve photon counting with a single stage

## Slide 9
Analogue Channel
Frame 1
Frame 2
Frame 3

## Slide 10
Analogue Channel
Frame 1
Frame 2
Frame 3

## Slide 11
Analogue Channel
Frame 1
Frame 2
Frame 3

## Slide 12
Analogue Channel
Amplifier
Folded cascode design
  High open loop gain
  Fast response time
Baseline voltage 400mV
  Generated at global level
Coarse stage has 3x higher bias current and different transistor sizing
  Increased response time to handle larger charge
Coarse stage has 3x higher bias current WHY

## Slide 13
Analogue Channel
Comparator
Differential pair with 2 gain stages
  Additional gain stage improves current balancing
  Improved response time
4-bit trim to minimise offset voltage
Coarse comparator sized for improved speed

## Slide 14
Analogue Channel
Pump time in multiples of Master Clock period
Current either pumped into channel or dumped to ground
Charge pump
Bias voltages set globally
  Calibrated to match number of photons
4-bit trim on pump current
  Trim circuit uses separate bias for increased flexibility
Maximum cancellation in course stage is 48 ph @ 25keV

## Slide 15
Pixel biasing
Bias voltages generated from DACs globally then buffered to channels
  DAC codes set via JTAG
  Set in initial calibration
Generated from 1mA current bandgap
GLOBAL DACS
COLUMN END
CHANNEL GROUP
CHANNEL
CHANNEL GROUP
CHANNEL
…
COLUMN END
CHANNEL GROUP
CHANNEL
CHANNEL GROUP
CHANNEL
…
…

## Slide 16
Analogue Channel
Configurable capacitance
Array of capacitors to suit operating mode
  Coarse Cf: 20 – 300 fF
  Transfer C: 80 - 320 fF
  Fine Cf: Fixed 50 fF
Power consumption
Each analogue channel consumes 220uA
  264 uW per channel
  7.3 W total

## Slide 17
Analogue Channel
Fine CSA output
Combined count
Coarse CSA output

## Slide 18
In-pixel RAM
Each pixel has a 256x16bit RAM
  Uses foundry provided 6T SRAM bit cell
  Full custom periphery
Enables increased functionality:
  Storing count values
  Accumulating values across frames
  Histograms
Estimated RAM will contribute ~60% of digital power for the pixel
  16-bit register available to bypass RAM

## Slide 19
Digital Control
Telegram
Operational commands sent through the external Telegram Interface on the fly
  Determines what is done with the pixel value
  Begins readout from pixels
Simple commands can be generated on-chip
Telegram encoded in 8b10b
Clocked at 68.45 MHz and synchronised to external interface

## Slide 20
Digital Control
Sequencer
Dynamic digital control signals generated by Sequencer on chip
  Configured through JTAG at start up
Controls:
  Reset
  Charge pump enable
  Charge transfer from Coarse to Fine
Clocked at 684.5 MHz and synchronised to start of Telegram Commands

## Slide 21
Long column mitigations
Maximised power tracks up columns for minimal IR drop (on M9 and M8)
Mirroring bias currents and buffer voltages per column
Channel simulated at 0.96 V for IR drop
  Accounting for a 120 mV drop along VDD and GND tracks
Sequencer signals generated per two columns to meet timing
Top pixels can be turned off

## Slide 22
Readout
Pixel values are shifted down columns into global logic at the bottom
Option for 2x2 pixel binning
Readout has up to 6 14.1Gbps serializer links available
  Used on previous test structures
  Can operate with 1, 2, 3, or 6 active
All serializers active
Left-most serializer active

## Slide 23
Design flow
Digital-on-top approach
  Pixels built in super pixel Channel Groups
  Channel Groups form Double Columns
  Double Columns form Pixel Matrix
  Global Circuitry added to Pixel Matrix
16 Analogue Channels
Digital logic
16 RAMs
(256x16b)

## Slide 24
Detector System
CZT detector bump bonded onto ASIC
Ribbon bonds to connect to PCB
  PCB folded underneath ASIC using jig
Full detector system will tile 4 ASICs together
  Expected to reach 85-90% fill of image area
Utilising existing STFC LOKI control boards
LOKI Control Boards

## Slide 25
Summary
Tape-out scheduled for end of June 2026
  Initial X-ray tests scheduled for Winter 2026
  Beamtime characterisation to follow
Test detector system in development with plans to scale up to Quad and 1 MP systems
Calibration
  Laser drilled tungsten mask
  Software method removing charge sharing events

## Slide 26
Notes: The abstract pattern can be removed or repositioned if required. Be careful to ‘Send to Back’ so that it does not obscure any important information.

## Slide 27
@STFC_matters
Science and Technology Facilities Council
Science and Technology Facilities Council
Notes: The abstract pattern can be removed or repositioned if required. Be careful to ‘Send to Back’ so that it does not obscure any important information.

## Slide 28
Backup

## Slide 29: Simple Software Extraction
Low Flux < 1 ph/frame
if neighbours > threshold charge sharing discard
else 0 or 1 photon in
central pixel  keep
Raw 500k frames
Decreasing threshold
1% of events
Still some charge loss where E lost < Th
Some correction to do?
PUBLICATION IN PROGRESS

## Slide 30: Physical Mask
20 µm
40 µm
Laser drilled array of holes on 110 µm pitch (16x16 trial)
Known Bad Pixel
No trim on test ASIC
40 µm
20 µm
Spacing of holes most important
Holes just need to be small enough
Works fine…
Probably not practical for deployment
But maybe a way to get correction for software method?
PUBLICATION IN PROGRESS