## Slide 1
Comparison of the simulated response of hybrid pixel detector with synchrotron beam tests
AG GDR, Mi2b
Marseille, le 29.06.2026
Mélissa Leroy1,2, Mathieu Dupont1, Christian Morel1, Elena Gaborieau-Borissenko2, Yannick Boursier1
1 Aix Marseille Univ, CNRS/IN2P3, CPPM, Marseille, France
2 Detection Technology, 38430 Moirans, France
melissa.leroy@deetee.com
mleroy@cppm.in2p3.fr

## Slide 2: Outline
1 – Context and goals of the thesis
2 – Modelling hybrid pixel detectors response
3 – Simulation pipeline
4 – Validation cases
5 – Conclusion and prospect

## Slide 3
Global project in Detection Technology (DT) company:
→ Develop Photon Counting Detectors (PCD) and associated technology solutions
→ Build an X-ray chain simulator as an internal research and development tool for the company
Motivation: why developing a simulation pipeline ?​
→ Optimize detector design ​
→ Generate training datasets for Deep Learning methods
→ Develop correction algorithms
→ Enhance physical understanding of the detector response
Thesis: ​​
Simulation of hybrid pixel Photon Counting Spectral Detectors
→ Detection Technology (DT) leverage the simulation expertise of imXgam team (CPPM)
1 - Context of the thesis
Medical X-ray imaging ​
Computed tomography​
Dental imaging​
Surgical imaging​
Security X-ray imaging​​
Baggage and parcel scan​​
Cargo and vehicle scan​​
People scan​​​​
Industrial X-ray imaging ​​​
Food industry​​​
Recycling and sorting​​​
Mining industry

## Slide 4
2 – Modelling hybrid pixel detectors response
  a. Ideal detector response
Ncharges : Number of electron-hole pairs created
Edeposited : Energy deposited in the sensor
Ee/h : Electron-hole pair creation energy
Ncharges
E e-/h
=
Material |  | Ee/h
Cadmium Telluride | CdTe | 4.43 eV
Cadmium-Zinc Telluride | CdZnTe | 4.64 eV
Silicon | Si | 3.64 eV
Germanium | Ge | 2.97 eV
Gallium Arsenide | GaAs | 4.2 eV
E deposited
Signal ∝ E γ

## Slide 5
2 – Modelling hybrid pixel detectors response
  b. Physical limitations of real detectors

## Slide 6
2 – Modelling hybrid pixel detectors response
  b. Physical limitations of real detectors
 |  | Atomic number | X-ray transitions energies |  | Mean free path (1/µ) | 
Material |  | Z | Kα1 [keV] | Kα2 [keV] | dα1 [µm] | dα2 [µm]
Si |  | 14 | 1.74 | 1.74 | 11.86 | 11.86
CdTe | Cd Te | 48 52 | 23.17 27.47 | 22.98 27.20 | 113.20 59.32 | 110.75 57.85

## Slide 7
2 – Modelling hybrid pixel detectors response
  b. Physical limitations of real detectors
N: Number of charges created
F: Fano factor
Material |  | Ee/h | Fano factor
Cadmium Telluride | CdTe | 4.43 eV | 0.24
Cadmium-Zinc Telluride | CdZnTe | 4.64 eV | 0.14
Silicon | Si | 3.64 eV | 0.115

## Slide 8
2 – Modelling hybrid pixel detectors response
  b. Physical limitations of real detectors

## Slide 9
2 – Modelling hybrid pixel detectors response
  b. Physical limitations of real detectors

## Slide 10
2 – Modelling hybrid pixel detectors response
  b. Physical limitations of real detectors
All these effect combine to shape the detector response (broadened peak, low-energy tail, escape peak)
To predict and disentangle them, we need a detailed physic-based simulation pipeline
→ And polarization, pile-up, threshold dispersion...

## Slide 11
Electronic noise added to the pixel collected charge spectrum reconstruction
New positions of charges after drift, diffusion and repulsion
Number of charges created and lifetime calculation
Geant4: Tracking of particles in the sensor, their position and energy loss
3 - Simulation pipeline
Output
Energy spectrum from the simulated detector
Counts in each of several predefined energy bins
Configuration
Detector geometry
X-ray source properties
Carrier generation and fluctuation
Charge carriers dynamics
Charge carriers
collection
  a. Global diagram
X-ray interactions
N: Number of charges
F: Fano factor
Ignores Shockley-Ramo signal induction, polarization,pile-up
σ(tdrift)
The 2 models studied :
Repulsion and diffusion are independent
E. Gatti et al., NIMA, 253 (1987) 393-399, DOI: 10.1016/0168-9002(87)90522-5.
Repulsion and diffusion are dependent
M.Benoit et al., NIMA, 606 (2009) 508-516, DOI: 10.1016/j.nima.2009.04.019.
Electrostatic repulsion
Diffusion

## Slide 12
CPPM: XPAD3.2-S/CdTe
2D Matrix 130 µm pitch 750 µm CdTe 1 threshold
Case 1: XPAD3.2-S/CdTe

## Slide 13
  a. SOLEIL acquisitions
4 - Validation case 1: XPAD3.2-S/CdTe
Synchrotron acquisition with XPAD3.2-S/CdTe
Threshold scan
Flat field 52 keV beam
A sample of 500 pixels reliable pixels
Calibration performed using two different spectral point
  Full energy peak
  Cd K-escape energy

## Slide 14
Source of discrepancies:
Higher noise than expected (close to 170 e – RMS)​
Non-uniform distribution of the threshold steps
Missing mechanisms from the simulation
  b. Simulation vs Measurements
4 - Validation case 1: XPAD3.2-S/CdTe
[3]: Repulsion and diffusion independents
[4]: Coupled repulsion and diffusion

## Slide 15
DT: prototype
2D Matrix 2x 24x36 pixels 350 µm pixel pitch 2 mm thickness CdTe / CZT 6 thresholds
Case 2: DT demonstrator

## Slide 16
4 - Validation case 2: DT demonstrator
Incident beam
z
y
z
x
y
Tile
z
x
y
ID11 source
10 x 10 µm² fine beam 59.5 keV / 74.9 keV
  a. ESRF ID11 experimental set-up
Beamtime (5 days allocated):
Characterize the prototype developed by DT and acquired experimental data to validate and refine detector simulation
Y. Boursier et al. (2026), “2D-CdZnTe detector prototype characterization for simulation validation,”Experimental report, Proposal MI-1525,Beamline ID11, ESRF, DOI: 10.15151/ESRF-ES-2266935027

## Slide 17
4 - Validation case 2: DT demonstrator
Heatmap of the matrix
Counts recorded for 3 adjacent pixels vs beam position beam position
Spectra of the target pixel and its neighbour
  b. Spot scans and spectral measurements

## Slide 18
4 - Validation case 2: DT demonstrator
  c. Spectral measurements vs simulation
Good agreement between simulated and measured spectra.
Accurate reproduction of the full-energy and fluorescence peaks.

## Slide 19
4 - Validation case 2: DT demonstrator
  d. Spot scans measurements vs simulation
SIMULATION
MESUREMENT
ONGOING WORK

## Slide 20
Goals of the study:
Simulation of high Z semiconductor pixel detectors response
Focus on the simulation of fluorescence escape and charge sharing
5 - Conclusion
Method:
Simulation integrates charge carriers' generation and dynamics
Study of 2 charge transport models, essential to reproduce accurately charge sharing
Processing of XPAD3.2-S/CdTe measurement performed in synchrotron
Measurement campaign spot beam performed with DT CdTe/CZT based prototype in Dec. 2025
Prospects:
Integration of signal induction on electrodes
Comparison with other simulation tools
Continue measurements analysis to refine the model
Results:
Good agreement
Behavior at this interpixel needs to be refined – ongoing work

## Slide 21: Thank you for your attention
Questions ?
Acknowledgement:
→ We acknowledge the European Synchrotron Radiation Facility (ESRF) for the provision of synchrotron radiation facilities under proposal ID MI-1525 and on beamline ID11. We thank Loic Jegou and the beamline staff for assistance and support during the beamtime.
→ We thank G. Gey, M.Plainemaison, T.Pierson, A. Winkler and M. Matikkala from Detection Technology for their support during data acquisition and experimental preparation.
AG GDR, Mi2b
Marseille, le 29.06.2026

## Slide 22
  b. X-ray interactions with Geant4
3 - Simulation pipeline
Tracking of particles in the sensor, their position and energy loss
Photoelectric effect, Compton scattering
CdTe : fluorescence escape photon
3 mm
1 mm
Americium 241
CdTe
3 mm
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection
Amptek : XR-100T-CdTe | 
 | Mono-pixel detector CdTe 3 mm pitch 1 mm thickness

## Slide 23
  c. Charge carriers dynamics
3 - Simulation pipeline
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection
Cathode
+
+
+
+
+
+
Anode
+
Charges drift toward the electrodes under the influence of electric field
Drift time is deduced from the position of the interaction z and a linear electric field

## Slide 24
2. Simulation framework
  c. Charge carrier dynamics
Diffusion
High concentration
Low concentration
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection
+
+
+
+
+
+
+
Repulsion
An inhomogeneous distribution induce a spontaneous Brownian motion
Charges of same sign repel each other due to electrostatic forces
Diffusion and repulsion cause a spontaneous spread of charge carriers
+
Electrons/Holes
Holes

## Slide 25
Equation solving
Complex + time consuming in simulation: approximation of the solution using models
The 2 models studied :
  Repulsion and diffusion are independent
E. Gatti et al., NIMA, 253 (1987) 393-399, DOI: 10.1016/0168-9002(87)90522-5.
- Repulsion and diffusion are dependent
M.Benoit et al., NIMA, 606 (2009) 508-516, DOI: 10.1016/j.nima.2009.04.019.
Diffusion and repulsion cause a spontaneous spread of charge carriers
Charge continuity equation
Describes the concentration of charge carrier evolution with respect to space and time
Neglecting generation and recombination at t > 0:
  c. Charge carriers dynamics
3 - Simulation framework
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection
Electrostatic repulsion
Diffusion
: Density of charges
: Diffusion coefficient
: Mobility of charge carriers
: Electric field induced by charge carriers

## Slide 26
  c. Charge carriers dynamics
2. Simulation framework
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection
Model 1: Decoupling repulsion and diffusion
Electrostatic repulsion
Diffusion
Repulsion-only equation:
Diffusion-only equation:

## Slide 27
  c. Charge carriers dynamics
2. Simulation framework
Model 1: Decoupling repulsion and diffusion
Diffusion-only equation:
For a spherical symmetry, the diffusion cause the electronic cloud to extend with
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection
Charges density follows 3D gaussian distribution
At t = 0, charges are concentered in a point

## Slide 28
  c. Charge carriers dynamics
2. Simulation framework
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection
Model 1: Decoupling repulsion and diffusion
Repulsion-only equation:
For a spherical symmetry, the repulsion cause the electronic cloud to extend and
Charges density follows a uniform distribution
At t = t0, charges are concentered in a point
Does not stand for energy deposition higher than 10 keV

## Slide 29
  c. Charge carriers dynamics
2. Simulation framework
Model 1: Decoupling repulsion and diffusion
For a cylindric symmetry, the repulsion cause the electronic cloud to extend and
Charges density follows a uniform distribution
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection

## Slide 30
  c. Charge carriers dynamics
2. Simulation framework
X-ray interactions
Carrier generation and recombination
Charge carriers dynamics
Charge carriers
collection
Model 2: Coupling repulsion and diffusion
Charges density follows 3D gaussian distribution
Electrostatic repulsion
Diffusion
At t = t0, charges are distributed

## Slide 31
4 - Validation case 1: XPAD3.2-S/CdTe
Simulation parameters
h mobility ( | 
h lifetime | 
Fano factor | 
Incident beam energy | 
Fano factor | 
Electronic noise approximated by a convolution with a gaussian
Fluorescence escape peaks are visible on the spectrum
The full energy peak is shifted due to an incomplete collection of charges
Activation of diffusion and repulsion increase the size of the electronic cloud and so the charge sharing
  b. Simulation

## Slide 32: ESRF set-up additional pictures

## Slide 33: ESRF set-up additional pictures

## Slide 34
4 - Validation case 2: DT demonstrator
  b. Spectral Detector Response - Simulation vs Measurements
Selection pixels with counts > 30

## Slide 35
Back-up. sigma vs energy: thresholds
Réponse spectrale d'un pixel pour un faisceau de 59.5 keV (droite) à et 79.5 keV (gauche)