## Slide 1
Réseau Semi-conducteur 11/06/26

## Slide 2: The integration of the DIAMMONI detector into ARRONAX and its application
Robin Molle
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 3: Why radiation tolerance ?
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Study of detector behavior in extreme conditions
Spatial
High Energy Physics
Medical applications

## Slide 4: Hadrontherapy application
Healthy tissues
Healthy tissues
Tumor
Conventional radiotherapy with X-rays
Kraan, A. C. (2015). Range verification methods in particle therapy: underlying physics and Monte Carlo modeling. Frontiers in oncology, 5, 150.
Dose = Energy deposited per unit mass
Radiotherapy: a method of cancer treatment
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Notes: + scientifique pour pic de bragg

## Slide 5: Hadrontherapy application
Kraan, A. C. (2015). Range verification methods in particle therapy: underlying physics and Monte Carlo modeling. Frontiers in oncology, 5, 150.
Hadrontherapy = Radiotherapy with ions
Dose = Energy deposited per unit mass
Radiotherapy: a method of cancer treatment
Healthy tissues
Healthy tissues
Tumor
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Notes: + scientifique pour pic de bragg

## Slide 6: Hadrontherapy application
Healthy tissues
Healthy tissues
Tumor
Kraan, A. C. (2015). Range verification methods in particle therapy: underlying physics and Monte Carlo modeling. Frontiers in oncology, 5, 150.
Online control in hadrontherapy
Dose = Energy deposited per unit mass
Radiotherapy: a method of cancer treatment
Hadrontherapy = Radiotherapy with ions
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Notes: + scientifique pour pic de bragg

## Slide 7: FLASH THERAPY: an innovative radiotherapy
Advantages of FLASH:
better preservation of healthy tissues for the same effectiveness at the tumor level
delivery of the entire dose in a fraction of a second
reduction in the number of sessions and treatment time
FLASH therapy
Ultra High Dose Rate
Conventional radiotherapy:
Dose rate: ~2 Gy/min = 0.03 Gy/s
FLASH therapy X :
Dose RATE: >10 kGy/s
FLASH therapy proton :
DOSE RATE: >100 Gy/s
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Gray (Gy) : energy per unit of mass
1 Gy = 1 J/kg

## Slide 8: DIAMMONI PROJECT CONTEXT
Online control in hadrontherapy
FLASH therapy
Ultra High Dose Rate
Diamond
Semi-conductor detector
DIAMMONI
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 9: DIAMMONI detector
DIAMMONI detector : a radiation detector for light ion beams at high flux and fluences (40 250 MeV)
Particle energy : 70 MeV (H-D-He)
Beam intensity
ARRONAX cyclotron :
ARRONAX AX casemate
ARRONAX beam
prototype for ARRONAX cyclotron
Beam structure :
Train duration : 50µs (« dt »)
Bunch duration : 4ns ON, ~29ns OFF
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 10
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Diamond
Semi-conductor detector
Flux response
Fluence response
Damaging ?
Loss of linearity ?
Flux relate to the speed of delivery
Fluence relate to the history of the device

## Slide 11: Solid State ionizing Chamber
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Diamond
+
+
-
Schokley-Ramo theorem :
Readout electronic
Applied to detection:
Ionizing particle
: weighting field
Ramo, S. (1939). Currents induced by electron motion. Proceedings of the IRE, 27(9), 584-585.
Shockley, W. (1938). Currents to conductors induced by a moving point charge. Journal of applied physics, 9(10), 635-636.

## Slide 12: Experimental setup
68 MeV proton range in diamond :
64 MeV alpha range in diamond :
68 MeV proton beam
64 MeV helium beam
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Fast oscilloscope
HV

## Slide 13: Main results in flux measurements
FLASH proton dose rate
1 kGy/s
Beam
Diamant
Protons 68 MeV
Alpha 64 MeV
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 14: Fluence measurements
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Charge collection efficiency:
Empirical model as a function of fluence
k: damage constant
d: thickness
: fluence
1st method:
Bäni, L., Alexopoulos, A., Artuso, M., Bachmair, F., Bartosik, M., Beck, H., ... & RD42 Collaboration. (2019). A study of the radiation tolerance of poly-crystalline and single-crystalline CVD diamond to 800 MeV and 24 GeV protons. Journal of Physics D: Applied Physics, 52(46), 465103.
Measure the k of diamond under ions
Irradiation
CCE measurement
CCE
2nd method :
Irradiation & CCE measurement
CCE calculated for each train

## Slide 15: Fluence experiments
Single crystal CVD Diamond
sCVD
Single crystal CVD Diamonds
1 diamond cut in 4
Helium
64 MeV
Proton
68 MeV
Spectroscopie alpha
XBIC - ESRF
Raman-PL
Recuit
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 16: Diamond response to Helium fluence
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 17: Diamond response to Helium fluence
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
17
XBIC map
Frenkel defects
Vacancies  NV centers
Annealing: 6h / 1000°C
64 MeV helium
irradiation
XBIC map
XBIC map

## Slide 18: Diamond response to Helium fluence
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
18
Annealing = increase of CCE
Frenkel defects created with irradiation
Electron and hole are not equivalent with irradiation
Vacancies (GR1 centers) trap charge carriers
GR1 converted with annealing in NV centers
After irradiation
Before annealing
After irradiation
After annealing

## Slide 19: Diamond response to proton fluence
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
M1 = control sample
Frenkel defects created with irradiation
Vacancies trap charge carriers

## Slide 20: Diamond response to proton fluence
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Electron and hole are not equivalent with irradiation
Annealing = increase of CCE

## Slide 21: Diamond response to proton fluence
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026
Excellent quality of the diamond
No particular impact on FWHM of Raman diamond peak
Vacancies converted with annealing in NV centers
H19 creation
NV centers creation depends on the substrate quality

## Slide 22: Model of CCE with fluence
Initial overall crystal quality
Trap filling
Heterogeneity of point defect distribution
Accumulation of irradiation-induced defects
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 23: To be published
Diamond and Related Materials

## Slide 24: DIAMMONI detector
High intensity mode
Integrated to ARRONAX AX beamline
Low intensity mode
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 25: DIAMMONI detector
High intensity mode
Low intensity mode
Two electronic boards developped for each mode:
Give the charge per train
Determine the real train duration
Reconstruct the time structure of the train
Give a spatial information on the beam
Give the average number of
particle per bunch
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 26: Take Home message
Study of diamond radiation tolerance as a function of fluence and flux to adapt it for various applications.
Demonstration of the linearity of the diamond response with flux over a wide range of dose rates.
Modelisation of the CCE with fluence
Material study of diamond after irradiation and annealing
DIAMMONI detector is integrated in ARRONAX, and was developed in ANR DIAMMONI project.
Prospects with helium ions, and also to see for a preclinical experiments using DIAMMONI detector calibration
Robin Molle – robin.molle@diamfab.com – Réseau Semi-conducteur 2026

## Slide 27
And now ?
  Start-up created in 2019
Based on more than 30 years of
R&D developed at Intstitut Néel- CNRS
+ 8 patents and know-how
  26 People
+ 5 advisors
Power electronics
Quantum application
Hardened electronics (space, nuclear and high temperature)
  Business targeted
Substrate
High Value-added Diamond Wafer
Diamond Electronic Devices
Diamfab unique know-how :
EPITAXY, DOPING, DEVICE DESIGN & PrOCESSING
GrenOBLE
FRANCE
IF you want to work with us, don’t hesitate to contact me !
robin.molle@diamfab.com
Robin Molle
Notes: So, at Diamfab we fabricate diamond.
Our solution is based on a technology developed for more than 30 years at CNRS Grenoble and protected with 8 patents and secrets know-how.
Today, Diamfab is 26 people and 5 world recognised advisors
We target power electronics and hardened electronics market as well as quantum market
Our unique know-how is to grow electronic grade doped diamond as a semiconductor on a diamond substrate and our expertise is to design and process devices.
Therefore, we have a dual model model allowing us to offer either High value-added wafer solution or bare die devices solution