## Slide 1: Toward a new gamma medical imaging technique: the XEMIS project
Presented by Samuel DUVAL
On behalf of the Xénon group
Industry-Academia Matching Event on Micro-Pattern Gaseous Detectors
26-27 April 2012, Annecy-le-Vieux

## Slide 2
Compton interactions could be also helpful for g-rays imaging
E1, x1, y1, z1
E2, x2, y2, z2
θ
Incident MeV g-rays
Good angular resolution (~ 1°) : good energy resolution ( ~ a few %)
good spatial resolution (< 0.5 mm in 3D)
Enough sensitive for g-rays imaging ? if depth of the detector is sufficient
With instrument containing
dense liquid (LXe) as detection medium
E0

## Slide 3
The XEMIS project
R&D focusing on MeV g-rays imaging with liquid xenon
and new radioisotopes produced by the ARRONAX cyclotron
The « gold » new prospect: 3g imaging with 44Sc
44Sc : detection of the 3rd 1.154 MeV g
in coincidence with the b+ decay
~1 cm resolution along the LOR achievable ?
C. Grignon et al., NIM A 571 (2007)

## Slide 4
XEMIS 1
“technical” proof of feasibility
Cryostat
Closed gaseous purification circuit
Liquefaction
tower
T. Haruyama team
PTR

## Slide 5
XEMIS 1 cryostat
Front-End Electronics
Insulating vacuum
High Volatge
Xenon outlet
Xenon inlet
MLI

## Slide 6
A LXe TPC with MICROMEGAS readout
Shaping rings
Segmented anode
6.35 x 6.35 mm2 pitch
12 cm
2.54 cm
Micromegas Grid
PMT

## Slide 7
TPC principle
Time of interaction:
Scintillation light (PMT)
T0 measurement
Energy and position:
Ionization with Micromegas and FEE
Drift time T1, E, x et y
Field rings
12 cm
300 m
Active volume
Recoil electrons create both scintillation and ionization in liquid xenon
PMT
Depth inferred from T0, T1 and the electron drift velocity :
z = (T1 – T0)·vdrift

## Slide 8
511 keV coincidence setup
22Na source
PMT
TPC
CsI Cristal
Collimator
FEE

## Slide 9
Resolution of interaction depth (@ 511 keV)
Profile of photoelectric interactions in the TPC
12 cm
Anode side
PMT side
Same method to measure the electron
drift velocity
300 µm measured
156 ns x 2 mm/µs
T. Oger et al. NIM A (2011)

## Slide 10
Monitoring of liquid xenon purity
λ = 19.52 ± 0.16 cm
S = 1.746 ± 0.006 V
S = 1.130 ± 0.004 V
Electronegative impurities absorbe drifting electrons
Z
S0 = 1.7623 ± 0.0015 V @511 keV

## Slide 11
Ionisation yield @ 511 keV
Good agreement with the Thomas model

## Slide 12
Measured and corrected energy resolution (511 keV)
T. Oger et al. NIM A (2011)
Very promising measured energy resolution

## Slide 13
XEMIS1 : what have we learned until now ?
Toward XEMIS2
“Technical” proof of concept:
    Intrinsic ionization energy resolution of liquid xenon is achievable with the low noise front-end electronics we developed
    Required spatial resolution for Compton reconstruction is achievable with Micromegas and liquid xenon
    Purification of an important liquid xenon volume is achievable by reproducing the methods used in fundamental researches (XENON100, ...)
What is missing to finish the phase I?
    Higher anode granularity in order to record the whole Compton sequence without unknown corners effect (2012 program)
What is missing to start the phase II?
    Industrial competences to extend the use of cryogenic liquids
    A wise choice of camera design

## Slide 14
XEMIS2 pre-design study for small animal imaging
with GATE1
Cylindrical camera XEMIS2 (~ 100 kg LXe)
Radius 8 < r < 20 cm
Length = 2×12 cm
Electric Field in z direction 2 kV/cm
192 PMTs (also for local trigger)
Micromegas ionization read-out
FEE Idef-X, pixels 3.175x3.175 mm2 (~25k channels)
Performances (simulation in progress, PhD student: A.F. Mohamad Hadi)
Efficiency to measure LOR: 30%
Efficiency to measure 1.157 MeV g-rays: 43%
3 photons efficiency (after selection): ~5%
Precision on localization along LOR ~ 1 cm (FWHM)
TPC characteristics
Intrinsic energy resolution: 5% @ 511 keV
Spatial resolution: 0.5 mm (X, Y and Z)
XEMIS2 LXe TPC
Rat Phantom
Cathode
Anode
Anode
PMTs
44Sc
Simulation status:
LXe Compton Telescope already implemented in GATE
Future => Simulation with test phantoms (NEMA, Derenzo…) with XEMIS2
1OpenGATE collaboration: http://www.opengatecollaboration.org/

## Slide 15
Expected to run at Subatech in 2014-15
Nantes Hospital in 2015-16
Rat phantom
PMTs
FEE for ionisation
read-out
Funding issues almost adressed
with the ARRONAXPLUS EQUIPEX
But some unknowns to solve :
- DAQ cost
- PMTs cost
511 keV LOR
3rd g
Improved reliability and safety :
ReStoX (liquid xenon station)
E
E
R&D on photodetection as an alternative of vacuum PMTs

## Slide 16
Gaseous Photomultiplier principle
Collection (THGEM)
Eampl
Ecoll ~ 0
Etrans
Eind
Entrance window (MgF2 ou SiO2)
Liquid xenon
Cathode
Anode
Ne mixtures
Energy loss
Transfer gap
Induction
(MICROMEGAS)
Amplification (PIM)
Transfer gap
Eampl
Eampl
500 lpi
670 lpi
CERN grid
UV photons
Photoelectric effect
THGEM
Reflective CsI
photocathode
THGEM: Efficient photoelectron collection + low gain
PIM/MICROMEGAS: ion blocking (prevents CsI damage) & gain
S. Duval et al., 2009 JINST 4 P12008 & S. Duval et al., 2011 JINST 6 P04007

## Slide 17
THGEM + supports
500 lpi grid*
670 lpi grid*
Kapton spacer (125 microns)
MICROMEGAS (50 microns)
Anode (ROGERS)
Base (stainless steel)
LXe GPM prototype (LXe side)
Vacuum side
Gas
SHV
THGEM/PIM/MICROMEGAS
Internal structure
Viewport (MgF2)
GPM detector

## Slide 18
Experiments in single-phase liquid-xenon
PGPM = 1100 mbar, T = 171 K, PXe = 1200 mbar, flow rate < 2 ln/h, DTin/out ~ 2K
Gas flow
meters
Pressure regulator
Gas outlet
Pump
Getter
PTR
Ne
CF4
Liquid xenon
Getter
GPM

## Slide 19
S. Duval et al., 2011 JINST 6 P04007
LXe Scintillation pulses recording
Ne/CF4 (90:10)
T = 173K
P = 1100mbar
GPM pulse
Source 238Pu
PMT Hamamatsu
(R7600-06MOD-ASSY)
GPM viewport
GPM
PMT
First pulse of a GPM in LXe !
Vacuum PMT pulse
GPM
Coincidence assembly

## Slide 20
S. Duval et al, NIM A (2011), doi: 10.1016/j.nima.2011.11.018
Gain measurements with 55Fe
55Fe source
Eextr1
Eextr2
Ecoll
Total gain above 106!

## Slide 21
55Fe Hybrid GPM pulse
25ns
Really like a PMT!
GPM
S. Duval et al., NIM A (2011), doi: 10.1016/j.nima.2011.11.018
Fast signal direct readout
50 Ohms

## Slide 22
5” window
Toward 20 inches diameter window or cylindrical one …
Proof of concept (2010)
High gains at LXe T ~106 (2011)
Efficient ion blocking expected also in Ne-mixtures
CsI photocathode studies in progress*
Large-size prototype is designed and ready for being assembled
Bulk-MICROMEGAS will be tested for future large-size cryogenic applications : SPLAM ANR
*A. Breskin et al., NIMA 639 (2011) 117-120 & S. Duval et al, NIM A (2011), doi: 10.1016/j.nima.2011.11.018
Conclusion & Prospects