## Slide 1: A Status Update on Proton Imaging for Applications in Medicine
Reinhard W. Schulte
Loma Linda University
pCT Collaboration
2012 NSS/MIC/RTSD Triple Joint Session

## Slide 2: Outline
Why should we image with protons rather than x-rays?
A short history of charged particle CT
Technical approaches to proton: single-particle detection, concepts and designs
Proton CT reconstruction
The quest for low-dose (sub-mSv) CT – is pCT the answer?

## Slide 3: Why should we image with protons?

## Slide 4: Tomographic Imaging Modalities for Proton Therapy: Primary & Secondary Use
X-ray CT: treatment planning, dose and range calculations, target volume delineation
MRI: target delineation, alignment verification
Positron Emission Tomography (PET)†: target volume definition, biologically weighted RT, treatment verification
MV/kV Cone Beam CT: alignment verification, adaptive radiotherapy

## Slide 5: The role of CT Imaging in Charged Particle Therapy
CT imaging is needed for
  Target volume definition (anatomical boundaries with additional information from fused MRI and PET studies
  Dose and range calculation
  Patient alignment verification (in-room cone beam CT, CBCT)

## Slide 6: Pros and Cons of X-ray CT
Pros
  Widely available (diagnostic CT has firm industry support)
  Thoroughly investigated and established
  Excellent image quality (spatial and density resolution)
Cons
  Issues with conversion of CT units to relative proton stopping power, leading to systematic range errors of the order of 3-5% for soft tissues and higher for tissues with very low or high density (lung, bone) or in the presence of metal artifacts
  Relatively high patient dose with CBCT when used as daily imaging modality for daily image verification
  Reconstruction artifacts with cone beam CT -> not suitable for dose replanning

## Slide 7: Range Accuracy of Proton/Ion Therapy when X-ray CT is used
X-ray attenuation in the 70-120kV range is determined by photo-electric and Compton scattering
Conversion of Hounsfield units to relative stopping power (RSP) requires careful calibration
There is no consistent one-to-one relationship between Hounsfield units and RSP for different tissues and materials
O. Jäkel, Imaging and Tumor Localization in Ion Beam Therapy, in: Ute Linz (Ed.), Ion Beam Therapy, Springer 2012)

## Slide 8: Process of X-ray CT to Relative Stopping Power Conversion
CT Hounsfield values (H)
Isodose distribution
Calibration curve
H = 1000 mtissue /mwater
Relative proton stopping power (RSP)
RSP = dE/dxtissue /dE/dxwater
Dose & range calculation

## Slide 9: Proton Beam Delivery Technique

## Slide 10: Charged Particle Treatment Planning: Proton Ray Tracing
Protons beams are modulated in energy such that their water-equivalent pathlength in tissue places the Bragg peak at the desired location
The most critical location is the “distal edge”
X-ray and pCT provide the required WEPL data

## Slide 11: Proton Ray Tracing Example (LLUMC)

## Slide 12: Motivation for pCT Range Uncertainties in Proton Therapy
Gensheimer F, et al., Int J Radiation Oncol Biol Phys 2010 (50% idosdose, red = observed, blue = planned
Differences in the interaction of x-rays and protons with matter make proton range calculations uncertain
Range uncertainties can range from mm to cm
Materials of unknown stopping power and CT artifacts create additional uncertainties
Proton CT is a potential solution to reduce this uncertainty from 3%-4% to ≤ 1% of range

## Slide 13: Goal: Reduction of Range Errors in Particle Therapy
Our ultimate goal in proton and ion therapy is to use proton and carbon ion beams for tomographic imaging.
This will allow reconstruction of the volumetric distribution of the proton ion stopping power relative to water, which is the information needed for proton and ion treatment planning.
First systems for proton imaging have been developed by us and others in recent years

## Slide 14: A short history of Charged Particle radiography and CT

## Slide 15: History of Proton Radiography (pRad)
A. Koehler was the first to point out the potential value of pRad and to perform experiments with 160 MeV (Koehler, Science 160, 303–304, 1968)
The higher density resolution but poorer spatial resolution than with x-ray radiography was noted by Koehler and later by Kramer et al. (Radiology, 1980)
Medical interest in pRad as a QA tool for proton therapy was revived by U. Schneider and E. Pedroni at PSA during the 1990s
pRad has been explored material testing at Los Alamos NL
pRad continues to be explored for clinical use at MGH (J Seco, this meeting) and PSI
More recently heavy ion radiography is being explored at the Heidelberg Ion Beam Therapy Center (K Parodi, this meeting)
Andy Koehler, former director of the Harvard Cyclotron)

## Slide 16: Proton Radiography as a Tool to measure Range Uncertainties
Alderson Head Phantom
Range Uncertainties
(measured with PTR)
> 5 mm
> 10 mm
> 15 mm
Schneider U. & Pedroni E. (1995), “Proton radiography as a tool for quality control in proton therapy,” Med Phys. 22, 353.

## Slide 17: History of Proton CT (pCT)
pCT is conceptually similar to pRad but consists of multiple pRad projections covering 360 deg followed by 3D image reconstruction
Proton CT was originally suggested by A. Cormack in 1963
First experiments were performed at LBL and LANL during the 1970s, and in Japan during the 1980s
Clinical proton CT as a low-dose tomographic imaging modality was explored by R. Martin at Argonne NL during the 1990s
A first pCT system was built by P. Zygmanski at the Harvard Cyclotron Lab in the late 1990s
A pCT collaboration between researchers interested in pCT was formed at BNL in 2003
pCT projects now exist at LLU/UCSC/CSUSB, NIU and FNAL, at INFN and IPHC

## Slide 18: Particle CT in the 1970s
Heavy ion tomography at LBNL 1972 – 1980 was explored by Cornelius Tobias and colleagues
Proton computed tomography was investigated at LANL by Ken Hanson et al.

## Slide 19: Proton Computed Tomography (pCT): Alan Cormack (Harvard, 1963)
Alan M. Cormack, physicist (1924-1998) was the first to publish a paper on the reconstruction of tomographic images based on X-ray absorption and proton degradation (J. Appl. Phys. 34, 2722, 1963)
It took less than 10 years before his idea became reality when the first when Godfrey Hounsfield constructed the first X-ray CT scanner
Both shared the Nobel Prize for Medicine in 1979
Alan M. Cormack, 1924-1998 Physics Nobel Laureate 1979

## Slide 20: Principles and Approaches

## Slide 21: Proton CT with Single Particle Detection – Evolution of Concepts
First concepts were derived from proton radiography systems (RWS talk Brazil 2001)
Fully developed concept of pCT (IEEE NSS/MIC 2003)
First pCT concept realized in 2010

## Slide 22: Single Particle Concept: Advantages and Challenges
Single particle detection allows for
  rejection of unsuitable events (“data cuts”)
  estimation of individual proton paths
  use of iterative reconstruction algorithms based on single proton histories
Challenges of single particle detection
  Requires high data rates (fast DAQ systems)
  Need to develop computation tools exploiting sparsity and massive parallelism

## Slide 23: Summary of pCT Principles
Detect location of individual primary particles, reject non related events (e.g., cosmic rays, noise)
Reconstruct charged particle path as accurately as possible
Measure residual energy or range of charged particle as accurately as possible
Convert residual energy/range to water-equivalent path length (WEPL)
Define the cpCT reconstruction problem for RSP using as much information as possible (mathematics-related)
Solve the cpCT reconstruction problem as efficiently as possible (computing hardware related)

## Slide 24: Current Approaches to Particle Tracking & Energy/Range Detection
Collaboration | Tracker | Energy/Range Detector
INFN | SSD | Crystal + PD
LLU/UCSC/NIU | SSD | Crystal + PD
NIU/FNAL | SciFi+SiPM | Range+WLSF+SiPM
LLU/UCSC/CSUSB | SSD | MSS (Plastic Scint) + PMT
TERA | GEM | Range+WLSF+SiPM
SSD – Silicon strip detector
Sci Fi – Scintillating fiber
SiPM Silicon photo multiplier
GEM Gaseous electron multiplier
PD – Silicon photodiode
WLSF – Wavelength-shifting fiber
MSS – Multi-stage scintillator
Hartmut F. W. Sadrozinski, et al. Detector Development for pCT, IEEE NSS/MIC 2011

## Slide 25: Why did we choose Silicon Strip Detectors?
Near 100% efficiency for particle detection with essentially zero noise occupancy
Inherently fine spatial resolution, no limitations as track estimation improves
Simple calibration that is stable over time for period of many years
Compact assembly using standard industry processes, with excellent mechanical stability
Robert P. Johnson, M09-78, IEEE NSS/MIC 2012

## Slide 26: LLU/UCSC/NIU pCT Tracker
The LLU/UCSC/NIU pCT tracker consists of front and rear module for location and direction measurements
Modules: two detector boards measuring the X-Y position in two locations => direction
Detector boards: 4 Si Strip Detectors (SSDs), 9 cm x 9 cm, 384 strips, 0.23 mm pitch
Strips oriented in horizontal or vertical direction (X and Y sensitivity)
Total sensitive area 9 cm x 18 cm
Modified GLAST/Fermi readout chip, max rate 200 kHz
Trackers/ASICs were developed/modified/tested by researchers & students from SCIPP
Phase I pCT scanner front module
Detector board with 2 SSDs in the front (visible) and 2 SSDs in the back of the board

## Slide 27: Problems with Current Tracker Design
Tiling of SSD leads to artifacts due to gap (a) or overlap (b) between/of sensitive areas
Confirmed by experimental scans

## Slide 28: Silicon Tracker Improvements
Large-area coverage requires tiling of Si sensors
Larges available sensors are ~9 cm x 9 cm single-sided Si strip detectors cut from 6 inch wafers (Hamamatsu)
Strip pitch 0.23 mm (shown not to be critical up to 1 mm by design study at NIU)
These sensors, in their native form have ~1mm inactive edges which would create image artifacts (see next slide for solution)
Layout of one of the four x-y tracker planes using 9 cm x 9 cm single-sided silicon detectors with one side rotated by 90 degrees.

## Slide 29: Si Sensors with “Slim” Edges
The corner of a sensor manufactured by Hamamatsu Photonics for the GLAST mission is shown on the left with the planned cut indicated by the red line
The sensor was etch-scribed with XeF2, cleaved and passivated with nitrogen plasma-enhanced chemical vapor deposition (PECVD)
The slim edge with strips and bias ring is shown on the right; only traces of the guard ring are visible.
Courtesy Hartmut F-W Sadrozinski

## Slide 30: Next Generation pCT Tracker/ASIC
4 (x/y) SSDs (9 cm x 9 cm) per tracking layer cut from 6” high-resistivity n-intrinsic silicon wafers
Eight layers of silicon strip detectors are read out by 144 64-channel ICs, each with a 100 Mbit/s link to an FPGA on the same board
Nominal data rate ~10^6 protons per sec
Poster M09-78 presented by Robert P. Johnson

## Slide 31: Next Generation pCT Data Readout/Data Flow
Readout is triggered by energy detector and tracker
Buffering at front end to minimize dead time
Eight layers of SSDs read out by 144 ASICs and linked to an FPGA (Spartan 6, Xilinx)
Poster M09-78 presented by Robert P. Johnson

## Slide 32: LLU/UCSC/NIU Crystal Calorimeter
Crystal matrix with 18 thallium-doped cesium-iodide (CsI(Tl)) crystals (~3.6 cm x 3.6 cm x 12.5 cm)
Each crystal read out by area-matched Si photodiode
Si photodiode => preamp/shaper => ADC
Excellent linearity and energy resolution < 1% above 40 MeV
Integrated with rear tracker module
Developed by V Bashkirov, V Rykalin & students from UCSC

## Slide 33: pCT Calorimeter Calibration
Careful calibration to convert weighted sum of crystal responses to WEPL
Performed prior to each pCT scan with polystyrene pates of known water-equivalent thickness (WET)
WET vs. mean calorimeter response fitted with 2nd degree polynomial to obtain WEPL calibration curve
Poster N14-189 presented by Ford Hurley

## Slide 34: Problems with Current Calorimeter
Despite careful calibration, inhomogeneous response to proton histories of same residual energy (WEPL) could not be resolved
Analysis revealed that this complicated response is due to light leakage at the level of the PDs and jitter of the timing of different front end chips

## Slide 35: Next Generation pCT Energy Detectors
Range counter consisting of stack of scintillator plates with SiPM readout and 3-5-stage plastic scintillator design (MSS) with PMT readout have been tested
Range counter determines residual range by detecting most distal plate with measurable response, WEPL = max range – residual range
MSS detects energy response of stopping scintillator, which is converted to WEPL
Both designs have comparable WEPL resolution (~3 mm) combined with good light yield and uniformity
MSS design preferred solution based on cost (including manpower) considerations
M22-21 Poster presented by Vladimir Bashkirov
N14-192 Poster presented by Andriy Zatserklyaniy

## Slide 36: The Phase 2 LLU/UCSC/CSUSB pCT Scanner
Large area tracker suitable for head scans + immobilization devices (36 cm x 9 cm
Multi (5)-stage scintillation detector calibrated for WEPL measurements
Vertical axis rotation stage for phantom rotation on fixed horizontal beam line
FPGA-based readout
Reconstruction on GPGPU cluster

## Slide 37: Phase 2 NIU/FNAL pCT Scanner
Scintillating fiber tracker
96-stage plastic scintillator range counter with WLSF read out by SiPMs (2 per stage)
Reconstruction on large GPGPU cluster (reconstruction of 1 billion histories in <10 min)
Courtesy George Coutrakon, NIU

## Slide 38: Proton CT image Reconstruction

## Slide 39: pCT Reconstruction Problem
Object is a “black box” containing a 3D distribution of target parameter (x) producing measured data y
Recovery of object function x requires inversion of physics model operator
Discrete linear problem approach requires x and y to be vectors and the physics operator to be represented by a matrix

## Slide 40: Linear pCT System and Solution Concept
Linear System
  i =1…m measurements => vector b = (bi)
  Object vector x = (xj), j=1…n
  Projection matrix A = (aij)
Note: matrix A is very large & sparse, and the system is inconsistent
Find “adequate” solution of Ax=b using iterative projection method (projection onto hyperplanes Hi)
Parallel algorithms for pCT have been developed by Yair Censor, Ran Davidi and Scott Penfold

## Slide 41: Path Reconstruction Concepts
Different proton paths may be used in the reconstruction: MLP = Most Likely Path, SLP = straight line path, CSP = cubic spline path
The MLP is determined by maximizing likelihood (chi square) of output parameters, given entry parameters
MLP significantly improves spatial resolution compared to SLP, CSP reconstruction is nearly as good as MLP reconstruction
The MLP concept was formulated by David C Williams (PMB 2004), & further refined by Scott Penfold, John Tafas and Keith Schubert (Med Phys 2008)

## Slide 42: First pCT Scanning Results with the LLU/UCSC/NIU pCT Scanner
A series of pCT scans of the Lucy radiosurgery phantom after completion and first calibration of the Phase I LLU-NIU-UCSC pCT scanner
Reconstructions were done with code written by graduate student Scott Penfold

## Slide 43: Complete Lucy Phantom Scan
Nov. 2011

## Slide 44: Proton Radiography
Cp radiographs give a much more faithful representation of relative bone and soft tissue electron densities
Experimental radiographs have been obtained with LLU/UCSC/NIU pCT prototype
pCT radiographs have proven useful in detecting flaws in current detector designs and data processing
Poster R04-24 presented by Tia Plautz

## Slide 45: PRIMA Collaboration
INFN pCT prototype, 4 layers of Si –trackers and YAG crystal calorimeter
Reconstructed image of a 4 cm x 2 cm PMMA cylinder with FBP using a Butterworth filter
Courtesy E Vanzi, INFN
M06-6 PRIMA Proton Imaging for Clinical Application. Presented by
Cinzia Talamonti.

## Slide 46: The quest of low-dose CT

## Slide 47: pCT – A low-dose image modality?
Dose per proton & image slice (180 x 180 x 1 mm3) of unit density: 0.67 x 0.5 nGy (<DE> = 100 MeV) from electronic interactions + 0.33 x 1 nGy (<DE> = 200 MeV) from inelastic nuclear interactions† = 0.67 nGy per proton history
Dimension of single slice head object vector: 1802 = 3.6 x 104 assuming 1 x 1 x 1 mm3 voxels
Number of protons per slice ~ number of unknowns x10-100 ~ 3.6 x 105 – 3.6 x 106
Average dose for a given slice: ~0.04-0.45 mSv (RBE = 1)
Comparable to 0.07 mSv from single chest X-ray
†Poster N14-186 presented by David Steinberg
R. Schulte, Med Phys. 2005

## Slide 48: Summary and Conclusions
pCT and radiography are evolving technologies that are driven by applications in charged particle therapy
Fastest progress is made by continued exchange of ideas and experience between medicine, high energy & nuclear physics, applied mathematics, computer science and engineering
Wide-spread use can be expected within the next 10 years as proton/ion therapy expands its role

## Slide 49: Acknowledgments
pCT research is funded by the National Institute of Biomedical Imaging and Bioengineering (NIBIB) and the National Science Foundation (NSF), award Number R01EB013118. The content of this paper is solely the responsibility of the authors and does not necessarily represent the official views of NIBIB and NIH. Work in pCT reconstruction is also supported by a grant from the United States - Israel Binational Science Foundation (BSF)
The proton imaging detectors were built at UCSC and Northern Illinois University with support from the U.S. Department of Defense Prostate Cancer Research Program, award No. W81XWH-12-1-0122and the Department of Radiation Medicine at LLUMC (Dr. James M. Slater)

## Slide 50: pCTlers
Vladimir Bashkirov
Ford Hurley
Andrew Wroe
Baron Black
Bob Jones
Hartmut Sadrozinski
Robert Johnson
Andriy Zatserklyaniy
Joel DeWitt
Forest Martinez-McKinney
Sergei Kashiguine
Edwin Spencer
Andrew Plumb
Cole Holcomb
David Steinberg
Vanessa Feng
Scott Macafee
Jessica Missaghian
Tia Plautz
Keith Schubert
Micah Witt
Blake Schultze
Reinhard Schulte
Scott McAllister
John Tafas
Yair Censor
Ran Davidi
Anatoly Rozenfeld
Peter Lazarakis
Valentina Giacometti
Carol Grande
Gerald Barnett
Gerald Blazey
George Coutrakon
Bela Erdelyi
Alexandre Dychkant
Eric Johnson
Sergey Uzunyan
Vishnu Zutshi
Karonis Nicholas
Kirk Duffin
Caesar Ordonez
John Winans
Andrew Gearhart
Mahmoud Yaqoub
Rick Salcido
Viktoriya Zvoda
Tony Zhang
Eric Olson
Michael Papka
Tom Uram
Peter Wilson
Greg Sellberg
John Rauch
John Krider
Paul Rubinov
Tom Fitzpatrick
Herman Cease
Michael Roman
Michael Utes
Anindya Sen
Gabor Herman
Scott Penfold
Márgio C. Loss Klock
Tianfang Li
Jerome Z. Liang
Ivan Evseev
Hugo Schelin
Olga Yevseyeva
Sergei Paschuk
Edney Milhoretto
Valeriy Denyak
David C. Williams
Todd Satogata
Steven Peggs
Klaus Mueller
Jason Heimann
Leah R. Johnson
Brian Keeney
David C. Williams,
Lan Zhang
Zheng Li
Craig Woody
Ron Martin
Ken Hansen
Feel free to join!
http://scipp.ucsc.edu/pCT/

## Slide 51: Thanks for listening!